## Settings and Navigation
Last Updated: Dec 18, 2020
Discussion: This setup is a simple "driving" or "Seating" using 2 pads and a Knee Switch and only supplies "Joystick Demands" to the wheelchair system. How the wheelchair reacts is entirely setp by the wheelchair electronics.
LEFT PAD |  | RIGHT PAD |  | CENTER PAD |  | USER PORT "Knee switch" |  | Patient's action | (Re) Action | Comments
Short Press | Long Press | Short Press | Long Press | Short Press | Long Press | Short Press | Long Press
None | LEFT | None | RIGHT | None | None | None | FORWARD | Patient activates left pad | Wheelchair gets LEFT demand for the duration of the left pad activation. | The wheelchair gets the demands because the "Short Press" settings are "None".
[8 empty] | Patient activates right pad | Wheelchair gets right demand for the duration of the right pad activation.
[8 empty] | Patient activates Knee Switch | Wheelchair gets forward demand for the duration of the Knee Switch activation.
Discussion: This setup uses 2 pads and a Knee Switch set to cycle between Forward and Reverse.
LEFT PAD |  | RIGHT PAD |  | CENTER PAD |  | USER PORT "Knee switch" |  | Patient's action | (Re) Action | Comments
Short Press | Long Press | Short Press | Long Press | Short Press | Long Press | Short Press | Long Press
None | LEFT | None | RIGHT | None | None | CYCLE | FORWARD, REVERSE | Patient activates left pad | Wheelchair gets LEFT demand for the duration of the left pad activation. | The wheelchair gets the demands because the "LEFT PAD Short Press" setting is "None".
[8 empty] | Patient activates right pad | Wheelchair gets right demand for the duration of the right pad activation. | The wheelchair gets the demands because the "RIGHT PAD Short Press" setting is "None".
[8 empty] | Patient activates and holds Knee Switch | Wheelchair gets forward demand for the duration of the Knee Switch activation. | The patients is holding the Knee Switch longer than the programmed "hold time" setting.
[8 empty] | Patient momentarily "taps" Knee Switch | The ASL110/165 swaps USER PORT setting from FORWARD to REVERSE. | See happens when the Knee Swtich is held… see next line.
[8 empty] | Patient activates and holds Knee Switch | Wheelchair gets reverse demand for the duration of the Knee Switch activation. | The patients is holding the Knee Switch longer than the programmed "hold time" setting.
Discussion: This setup uses 2 pads cycle between Left and Right and between Forward and Reverse.
LEFT PAD |  | RIGHT PAD |  | CENTER PAD |  | USER PORT "Knee switch" |  | Patient's action | (Re) Action | Comments
Short Press | Long Press | Short Press | Long Press | Short Press | Long Press | Short Press | Long Press
CYCLE | LEFT, RIGHT | CYCLE | FORWARD, REVERSE | None | None | None | None | Patient activates left pad | Wheelchair gets LEFT demand for the duration of the left pad activation.
[8 empty] | Patient momentarily "taps" Left Pad | The ASL110/165 swaps the Left Pad from LEFT to RIGHT.
[8 empty] | Patient activates left pad | Wheelchair gets RIGHT demand for the duration of the left pad activation.
[8 empty] | Patient momentarily "taps" Left Pad | The ASL110/165 swaps the Left Pad from RIGHT back to LEFT.
[8 empty] | Patient activates left pad | Wheelchair gets LEFT demand for the duration of the left pad activation.
[8 empty] | Patient activates right pad | Wheelchair gets FORWARD demand for the duration of the right pad activation.
[8 empty] | Patient momentarily "taps" Right Pad | The ASL110/165 swaps the Left Pad from FORWARD to REVERSE.
[8 empty] | Patient activates right pad | Wheelchair gets REVERSE demand for the duration of the right pad activation.
[8 empty] | Patient momentarily "taps" Right Pad | The ASL110/165 swaps the Left Pad from REVERSE to FORWARD.
Discussion: This setup uses 2 pads and a Knee Switch to cycle through programmed features.
LEFT PAD |  | RIGHT PAD |  | CENTER PAD |  | USER PORT "Knee switch" |  | Patient's action | (Re) Action | Comments
Short Press | Long Press | Short Press | Long Press | Short Press | Long Press | Short Press | Long Press
CYCLE | LEFT, FORWARD, REVERSE | CYCLE | RIGHT | None | None | DO FEATURE | NEXT FEATURE | Patient momentarily "taps" Left Pad | The ASL110/165 cycles from LEFT to FORWARD to REVERSE on each "tap". A long hold of the Left Pad executes the setting.
[8 empty] | Patient momentarily "taps" Knee Switch | The ASL110/165 executes the "ACTIVE FEATURE". Example: If POWER ON/OFF is active, a "tap" causes the ASL110/165 to "go to sleep". | This mimics the existing USER PORT operation.
[8 empty] | Patient holds the Knee Switch active. | The ASL110/165 changes to the next programmed feature, i.e. POWER ON/OFF, BLUETOOTH or NEXT MODE.
Discussion: This setup uses 3 pads and a Buddy Button to cycle through programmed features. This setup mimics the existing/standard ASL110/ASL165 operation.
LEFT PAD |  | RIGHT PAD |  | CENTER PAD |  | USER PORT "Knee switch" |  | Patient's action | (Re) Action | Comments
Short Press | Long Press | Short Press | Long Press | Short Press | Long Press | Short Press | Long Press
None | LEFT | None | RIGHT | None | FORWARD | DO FEATURE | NEXT FEATURE | Active a Pad.. It goes that way.
Discussion: This setup uses 2 pads and a Knee Switch to cycle through programmed features and FWD/REV.
LEFT PAD |  | RIGHT PAD |  | CENTER PAD |  | USER PORT "Knee switch" |  | Patient's action | (Re) Action | Comments
Short Press | Long Press | Short Press | Long Press | Short Press | Long Press | Short Press | Long Press
None | LEFT | None | RIGHT | None | None | CYCLE | Programmed Features, FORWARD, REVERSE | The patient taps the "Knee Switch". | The ASL110/165 activates one feature then the next, then the next, then the next, then to FWD, then to REVERSE then back again.
[8 empty] | The patient holds the "Knee Switch". | The featue is executed. - If POWER ON/OFF is active, the long hold turns the ASL110/165 on and off. - If BLUETOOTH is active, sends a forward to the Bluetooth module. - If NEXT FUNCTION is active, a short signal is activated on Pin 5. - If NEXT PROFILE is active, a longer signal is set at Pin 5. - If FORWARD is active, a long hold sends a Forward demand to the system. - If REVERSE is active, a long hold sends a Reverse demand to the system. NOTE: Each feature can be enabled/disabled.
Discussion: This setup will use the Left Pad to cycle through selections and the Right Pad to execute the selection. This is a two-switch solution. When the selection has been executed, the list is reset to the first selection.
LEFT PAD |  | RIGHT PAD |  | CENTER PAD |  | USER PORT "Knee switch" |  | Patient's action | (Re) Action | Comments
Short Press | Long Press | Short Press | Long Press | Short Press | Long Press | Short Press | Long Press
CYCLE RIGHT PAD | None | None | DO FEATURE | None | None | None | None | The patient taps the Left Pad.
CHOP: Don't forgot abut speed changing
CHOP: Changing templates/configurations during the day by the patient is a huge advantage.
(Added 8/21/20 Gchop) I think we can implement a "super long hold" or other button sequence to switch between User Profiles.
Note on E39: Gary Chopcinski: This might be an oops!

## Programming
LEFT PAD
Short Press | Long Press
None | Left
Cycle | Right
Do Feature | Forward
Reverse
| POWER ON/OFF |  | Profile Selection Screen: This is the first screen of programming.
| BLUETOOTH |  | Every programming screen would/could show the Diamod Profile number or No icon if it's a Global/System setting.
NEXT FUNCTION
NEXT PROFILE
TOGGLE F/R
USER MENU
SEATING
Add Sleep options
This screen shows the Pad Direction if the SHORT PESS is set to NONE.
Behaves as it does today.
Note that the CENTER PAD is set for CYCLING. The "Cycled" features are enabled in the LONG PRESS setting screen.
OR the following screen

## Newer Main Screen
This is "CHANGE TO NEXT GROUP OF SETTINGS" feature. | This is AUDIBLE OUT feature
Long USER PORT SWITCH press
Long USER PORT SWITCH press
Long USER PORT SWITCH press
Long USER PORT SWITCH press
Short
Short
Short
Long USER PORT SWITCH press (TYPICAL)
Long USER PORT SWITCH press

## User Screen
2 SWITCH
Note: Can be any combination of 2 switches (directions)
|  | 1 SWITCH |  |  |  | 3 SWITCH |  | 4 SWITCH
| Note: Can be any switch (direction) |  |  |  | Note: Can be any combination of 3 switches (directions) |  | Note: All 4 switches (directions) are used.
FUTURE
LEFT PAD: Set with SHORT PRESS: NONE
LONG PRESS: LEFT
MODE PORT: Set for Short Press for CYCLE. LONG PRESS has at least NEXT FUNCTION active.
USER PORT: Set for Short Press for CYCLE. LONG PRESS has at least POWER ON/OFF active.
RIGHT PAD: Set with SHORT PRESS: NONE
LONG PRESS: RIGHT
2 PAD's: Set with SHORT PRESS: NONE and LONG PRESS: NONE
LEFT PAD: Set with SHORT PRESS: NONE
LONG PRESS: LEFT
3 PAD's: Set with SHORT PRESS: NONE and LONG PRESS: NONE
This area depicts the Group/Profile annunciated using color.